Comparison of Manual and Automated Measurements of Reactive Gaseous Mercury (RGM)

Katrine Aspmo1*, Enno Bahlmann2, Cathy Banic3, Torunn Berg1, Alexandra Steffen3 & Christian Temme2

1 Norwegian Institute for Air Research, P.O.BOX 100, 2027 Kjeller, NORWAY, 
E-mail: katrine.aspmo@nilu.no

2 GKSS Research Centre, Geesthacht, GERMANY

3 Environment Canada (MSC), Toronto, CANADA

Abstract: A measurement campaign was undertaken in Ny-Ålesund, Spitsbergen (78°55’N, 11°90’E) during April and May 2003. Four mercury depletion events (MDE) were observed during the study. Three research groups measured reactive gaseous mercury (RGM) at three different measurement spots. All groups observed increased concentrations of RGM as concentrations of GEM decreased. Recorded concentrations of RGM ranged from less than 7 to 450 pg/m3. Dfferences between both measured parallels and the measurements sites were observed, even though no significant differences in the mean GEM concentrations were revealed between the measurement spots.

Key words: Mercury depletion events, MDE, Reactive gaseous mercury, RGM, comparison,

Introduction

Speciation of atmospheric mercury is important in order to understand the fate of mercury in the environment and mercury’s implications on the environment. For the past 5-10 years a lot of work have been undertaken to develop methods to measure inorganic atmospheric forms of mercury, referred to as reactive gaseous mercury (RGM). Different groups has studied RGM at various locations through polar regions, and to ensure comparable data sets an international field campaign was undertaken in Ny-Ålesund, Spitsbergen (78°55N, 11°90E), during April and May, 2003.

This work presents measured concentrations of reactive gaseous mercury (RGM) obtained by three different research groups at three different measurement sites. One group, NILU, measured at the Zeppelin station, which is located at the Zeppelin mountain, 474 m.a.s.l. The station is typically positioned well above frequent ground inversion of the fjord valley. The second group, MSC, measured at the Italian station, which is located in the Ny-Ålesund village, 12 m.a.s.l. The third group, GKSS, measured at “Ny-FID Sund”, which is located 500 m east of Ny-Ålesund at about 10 m.a.s.l. The groups used the same principle for collection and quantification of RGM as described by Landis et al, 2002. RGM was collected on KCl coated annular denuder tubes followed by thermal decomposition of RGM quantified as Hg0. Further details are shown in table 1.

Table 1: Summary of methods used by 3 research groups for measurements of RGM.

|  | NILU | MCS | GKSS |
| --- | --- | --- | --- |
| Sampling system | Manuel | Manuel | Tekran 1130 |
| Number of parallels | 2 | 2 | 1 |
| Sampling site (m.a.s.l.) | 474 | 12 | ~10 |
| Denuder length : length of active surface (cm) | 37,0 : 20,0 | 50,8 : 25,4 | 51,4 : 25,4 |
| Sampling flow (lpm) | 5 | 10 | 8,5 |
| Sampling temperature (°C) | 30 – 40 | 50 - 60 | 50 |
| Desorbtion system | Lindberg furnace | Lindberg furnace | Tekran 1130 |
| Desorbtion temperature (°C) | 500 | 500 | 500 |
| Quantification system | Tekran 2537A | Tekran 2537A | Tekran 2537A |

Results and discussion

Four MDE’s where detected during the field study, and they are labeled according to Steffen et al, 2004. During non-MDE situations all groups measured RGM-concentrations at the methods detection limit (MDL = 7 pg/m3). In general during all events all groups measured increased concentrations of RGM while concentration of GEM decreased, as shown in figure 1 (MDE 1 is not shown in the figure).

Figure 1: Development of RGM-concentrations during 3 MDE’s

During MDE 1 small concentrations of RGM were observed (RGM < 16 pg/m3) and the agreement between the sampling sites were good. During MDE 2a higher RGM was observed (max RGM = 80 pg/m3). The precision between sets of parallels were in general good (except for one sample), but not the agreement between the sites. Both sites where at sea level, but one site shows 3 times higher values than the other. Concentrations of RGM increased further during MDE 2b. The RGM patterns observed at sea level are similar, but off-set by a factor of 3, with the highest RGM-concentration observed when GEM started to increase. RGM observed at Zeppelin were highest earlier in the event when GEM concentrations were the lowest. The precision between the parallels seemed to decrease as the concentration of RGM increased. During MDE 3 all groups measured comparable concentrations of RGM (overall precision within 30% except for MSC – 2). The precision between parallels measured at Zeppelin was good, except for one sample (overall precision within 15%). The variability between parallels measured at sea level was high. The RGM-pattern observed during MDE 4 shows two peaks and the measured RGM-concentrations the highest correlation (r2 = 0.93). The agreement between the sites were relatively good, with a dicrepancy of 30%.

Even though the mean GEM concentrations reveals no significant difference between sea level and the Zeppelin station (Berg et al, 2003, Temme et al, 2004) we can not, with our current understanding, deduce that RGM measured at the same sites would show similar behavior. The authors concern is the huge differences between parallel samples. Preliminary results from Barrow, Alaska spring 2004 show similar behavior.

Of the two groups measuring manual RGM, MSC followed the guidelines set by Landis et al, 2002 whereas NILU made a few adjustments, such as reduced sampling flow and shorter denuder length. Both groups used impactors (cut off 2,5 µm) as sample inlet. The large differences between the parallel samples might be due to coarse particles entering the system and thus adsorbed to the denuder walls. This is a plausible explanation for the group using reduced sample flow. None of the above mentioned groups were able to attain comparable results. The Arctic is a harsh environment, which might be a possible explanation why comparable parallel samples were not obtained.

Conclusions

A comparative study of RGM was conducted in Ny-Ålesund, Spitsbergen during spring 2003. All groups observed increased concentrations of RGM as concentration of GEM decreased. The recorded concentrations of RGM differed between the sampling sites and also within sampled parallels. At this point we need more experience to understand whether the differences between the sites are due to real differences or not. The groups’ inability to produce comparable parallels reveals the fact that at this point the method is not robust enough to be credible.

Acknowledgements

This work is founded by the Research Council of Norway. Thanks to the Norwegian Polar Institute for technical, logistical and financial support during the field campaign, and also a special thanks to the Ny-Ålesund community for hospitality, friendliness and warmth.

## References

Berg T., Sommar J., Wangberg I. Gårdfeldt K., Munthe J. And Schroeder W.H. (2003) Arctic mercury depletion events at two elevations as observed at the Zeppelin station and Dirigible Italia, Ny-Ålesund, spring 2002, J. Phusics

Landis, M., Stevens, R.K., Schaedlich, F. & Prestbo, E.M. (2002): Development and characterization of an annular denuder methodology for the measurement of divalent inorganic reactive gaseous mercury in ambient air. Environmental Science and Technology; Vol. 36, pp. 3000-3009.

Steffen, A., Aspmo, K., Banic, C., Berg, T., Ebinghaus, R., Ferrari, C., Gauchard, P.-A., and Temme, C. (2004): Characterizing Mercury Depletion Events in Ny-Ålesund, 2003. Presented at the ICMGP in Ljubljana, Slovenia, June 27 – July 2, 2004.

Temme, C., Aspmo, K., Bahlmann, E., Banic, C., Berg, T., Dommergue, A., Ebinghaus, R., Ferrari, C., Gauchard, P.-A., Magand, O., Pirrone, N., Planchon, F., Sprovieri, F. and Steffen, A., (2004): Repeatability and reproducibility of data from different groups and locations in Ny-Ålesund during the Hg-campaign 2003. Presented at the ICMGP in Ljubljana, Slovenia, June 27 – July 2, 2004.